TD4C PRP/PRF/GFC Centrifuge.

General Operating Instructions

Important: These are general guidance notes. Always follow the protocol recommended by the manufacturer of the collection tubes and the treating physician.

Before You Start

Place the centrifuge on a stable, level surface.

Connect to the correct power supply (typically 220V/50Hz).

Inspect rotor and tubes for damage.

Always balance the rotor with tubes of equal weight placed opposite each other.

Close and lock the lid before starting.

General Operating Steps

Collect blood using the appropriate sterile collection tube.

Load balanced tubes into the rotor.

Close the lid securely.

Set the required RPM and time according to the tube protocol.

Press Start.

Wait until the rotor stops completely.

Open the lid and carefully remove the tubes.

Proceed with sample collection according to the treatment protocol.

Recommended General Settings

| Tube Type | Typical RPM | Typical Time | Purpose |
| --- | --- | --- | --- |
| PRP (ACD + Gel) | 3700–4000 RPM | 5–8 min | Platelet-Rich Plasma |
| PRF | 2700 RPM | 7–12 min | Platelet-Rich Fibrin |
| GFC | 3400–3600 RPM | 8–10 min | Growth Factor Concentrate |

Safety Instructions

Never open the lid while the rotor is spinning.

Do not overload the rotor.

Replace damaged tubes immediately.

Clean the chamber after use.

Disconnect power before maintenance.

Maintenance

Clean rotor after each working day.

Inspect rotor and lid regularly.

Store in a dry, dust-free environment.